## Sheet1
Test Equipment/Chemicals | Identification Numbers | Least Count and Range | Valid Calibration | Tests Used | Clause Reference | Remarks | Instrument Status | Calibration Status
Voltmeter | 1 | 0.1V, 0-600V | 19.05.22 | Power Input & Current | Cl.10 of IS 302-2-30 | Lab MTP 2 | Ekta | Required
Ammeter | 2 | 0.01A, 0-20A | 19.05.22 | Power Input & Current | Cl.10 of IS 302-2-30 | Lab MTP 2 | Ekta | Required
Wattmeter | 3 | 1W, 0-4000W | 19.05.22 | Power Input & Current | Cl.10 of IS 302-2-30 | Lab MTP 2 | Ekta | Required
IR Meter | 4 | NA,2 - ∞ MΩ | 19.05.22 | Leakage Current & Electric Strength at operating Temperature | Cl.13 of IS 302-2-30 | Lab MTP 2 | Ekta | Required
Micro ampere Meter | 5 | 1uA, 0-500uA | 19.05.22 | Leakage Current at operating Temperature | Cl.13 & 16 of IS 302-2-30 | Lab MTP 2 | Ekta | Required
Voltmeter with finger set | 6 | 0.01V, 99.99 V | 19.05.22 | Protection against access to live parts | Cl. 8 & 21of IS 302-2-30 | Lab MTP 2 | Ekta | Required
Dead weight | 7 | 5 kg | 19.05.22 | Protection against access to live parts | Cl. 8 & 21of IS 302-2-30 | Lab MTP 2 | Ekta | Required
Frequency Meter |  | 0.01Hz, 45-60Hz | 29.09.21 |  |  |  | OEGPL | Available
Stopwatch | 8 | 0.01Sec, 0-24Hr | 19.05.22 | Leakage Current & Electric Strength at operating Temperature | Cl.13 & 16 of IS 302-2-30 |  | OEGPL | Required
Digital kV meter | 9 | 0.001kV, 0-5kV | 29.09.21 | Electric Strength at operating Temperature | Cl.13 & 16 of IS 302-2-30 | HV Tester | OEGPL | Available
Mili Ampere Meter | 10 | 1 mAmp, 0-250mAmp | 29.09.21 | Electric Strength at operating Temperature | Cl.13 & 16 of IS 302-2-30 | HV Tester | OEGPL | Available
Voltmeter ECRT |  | 0.01V, 9.99 V | 19.05.22 | Provision for Earthing | Cl. 27 of IS 302-2-30 | Lab MTP 1 | OEGPL | Required
Ammeter ECRT |  | 0.1A, 0-30A | 19.05.22 | Provision for Earthing | Cl. 27 of IS 302-2-30 | Lab MTP 1 | OEGPL | Required
Dig Ampere Meter | 11 | 0.1A, 0-200A | 19.05.22 | Resistance to Heat & Fire | Cl. 30 of IS 302-2-30 | Glow wire Appratus | OEGPL | Required
Dig Temperature Controller | 12 | 1˚C, 0-1200˚C | 19.05.22 | Resistance to Heat & Fire | Cl. 30 of IS 302-2-30 | Glow wire Appratus | OEGPL | Required
Dig Timer Conroller | 13 | 0.1Sec, 0-99.9 Sec | 19.05.22 | Resistance to Heat & Fire | Cl. 30 of IS 302-2-30 | Glow wire Appratus | OEGPL | Required
Digital Temperature Meter with 5 Sensor Wire | 14 | 0.1˚C, 0-300˚C | 19.05.22 | Resistance to Heat & Fire | Cl. 30 of IS 302-2-30 |  | Ekta | Required
Ball Pressure Appratus | 15 |  | 19.05.22 | 1- Mechanical Strength 2- Resistance to Heat & Fire | Cl. 21 & 30 of IS 302-2-30 |  | Ekta | NA
Travelling Microscope | 16 | 0.01mm, 0-150mm | 19.05.22 | Resistance to Heat & Fire | Cl. 30 of IS 302-2-30 |  | OEGPL | Required
Dead Weight | 17 | 6 kg | 19.05.22 | Supply connection & External Flexible cords | Cl. 25 of IS 302-2-30 | Cord Grip Test Appratus | OEGPL | Required
Measuring Tape |  | 1mm, 0-1000 mm | 29.09.21 | Supply connection & External Flexible cords | Cl. 25 of IS 302-2-30 |  | OEGPL | Available
Micrometer |  | 0.001, 0-25mm | 29.09.21 | 1- Supply connection & External Flexible cords 2- Clearance, Creepage & Solid Insulation | Cl. 25 & 29 of IS 302-2-30 |  | OEGPL | Available
Filler Gauge |  | 0.03, 0.03 - 1mm | 29.09.21 | Clearance, Creepage & Solid Insulation | Cl. 29 of IS 302-2-30 |  | OEGPL | Available
Vernier Caliper |  | 0.02mm, 0-150mm | 29.09.21 | Clearance, Creepage & Solid Insulation | Cl. 29 of IS 302-2-30 |  | OEGPL | Available
Torque Screw Driver | 20 | 0.2 nm,0.1-1.5 nm | 19.05.22 | 1- Supply connection & External Flexible cords 2- Screws & Connection | Cl. 25 & 28 of IS 302-2-30 |  | Ekta | Required
Test Corner, Felt Strips, Copper/Brass disc | 21 |  |  | Heating | Cl. 11 of IS 302-2-30 |  | Ekta | NA
Dig Temperature Controller | 22 | 0.1˚C, 0-300˚C | 19.05.22 | Heating | Cl. 11 of IS 302-2-30 | Electric Hot air Oven | OEGPL | Required
Digital Hour Meter | 23 | 0.01Hr, 0-99999.99Hr | 19.05.22 | Heating | Cl. 11 of IS 302-2-30 | Electric Hot air Oven | OEGPL | Required
Inclined Plane | 24 |  |  | 1- Mechanical Strength 2- Resistance to Heat & Fire | Cl. 21 & 30 of IS 302-2-30 |  | Ekta | Required
Spring Hammer | 25 | 0.5j, 0-5 J | 19.05.22 | 1- Mechanical Strength 2- Resistance to Heat & Fire | Cl. 21 & 30 of IS 302-2-30 |  | Ekta | Required
Dig. Temp Controller | 26 | 0.1˚C, 0-99.9˚C | 19.05.22 | Moisture Resistance / Resistance to Rusting | Cl. 15 & 31 of IS 302-2-30 | Humidity Chamber | OEGPL | Required
Dig. Hygrometer | 27 | 1%Rh, 0-99%Rh | 19.05.22 | Moisture Resistance / Resistance to Rusting | Cl. 15 & 31 of IS 302-2-30 | Humidity Chamber | OEGPL | Required
Hour Meter | 28 | 0.01Hr, 0-99999.99Hr | 19.05.22 | Moisture Resistance / Resistance to Rusting | Cl. 15 & 31 of IS 302-2-30 | Humidity Chamber | OEGPL | Required
Mili Ohm Meter | 29 | 0-19.9 ohm, Variable | 19.05.22 | Supply connection & External Flexible cords | Cl. 22, 25, 29 of IS 302-2-30 |  | OEGPL | Required
Angle D | 30 | 0.5˚, 180˚ | 19.05.22 | Stability & Mechanical hazard | Cl. 20 of IS 302-2-30 | Inclined Test Appratus | Ekta | Required
Digital Voltmeter | 31 | 0.1V, 0-500V | 19.05.22 | Endurance Test | Cl. 18 of IS 302-2-30 | Endurance Test Panel | Ekta | Required
Digital Ampere meter | 32 | 0.01A, 0-8A | 19.05.22 | Endurance Test | Cl. 18 of IS 302-2-30 | Endurance Test Panel | Ekta | Required
Digital Counter meter | 33 | 1C, 999999C | 19.05.22 | Endurance Test | Cl. 18 of IS 302-2-30 | Endurance Test Panel | Ekta | Required
Digital timer cycle On/Off | 34 | 1min, 999 min | 19.05.22 | Endurance Test | Cl. 18 of IS 302-2-30 | Endurance Test Panel | Ekta | Required
Digital Watt Meter | 35 | 1W, 2000W | 19.05.22 | Endurance Test | Cl. 18 of IS 302-2-30 | Endurance Test Panel | Ekta | Required
Hour Meter | 36 | 0.01Hr, 0-9999.99Hr | 19.05.22 | Endurance Test | Cl. 18 of IS 302-2-30 | Endurance Test Panel | Ekta | Required
Torque Wrench | 37 | 0.5nm, 3-14nm | 19.05.22 | Screw & Connection | Cl. 28 of IS 302-2-30 |  | Ekta | Required
Digital VAW Meter |  | 0.01V, 0.1W, 0.001A/ 0-600V, 0-15000W, 0-25A | 29.09.21 | Power Input & Current | Cl.10 of IS 302-2-30 | Line Testing Panel | OEGPL | Available
Micro ampere Meter |  | 0.1uA, 0-2000uA | 29.09.21 | Leakage Current at operating Temperature | Cl.13 of IS 302-2-30 | Line Testing Panel | OEGPL | Available
IR Meter |  | NA,2 - ∞ MΩ | 29.09.21 |  | Cl.13 of IS 302-2-30 | Line Testing Panel | OEGPL | Available
Voltmeter ECRT | 38 | 0.01V, 9.99 V | 29.09.21 | Provision for Earthing | Cl. 27 of IS 302-2-30 | Line ECR Tester | OEGPL | Available
Ammeter ECRT | 39 | 0.1A, 0-30A | 29.09.21 | Provision for Earthing | Cl. 27 of IS 302-2-30 | Line ECR Tester | OEGPL | Available
Digital kV meter | 9 | 0.001kV, 0-5kV | 29.09.21 | Electric Strength at operating Temperature | Cl.13 & 16 of IS 302-2-30 | HV Tester Line Testing | OEGPL | Available
Mili Ampere Meter | 10 | 1 mAmp, 0-250mAmp | 29.09.21 | Electric Strength at operating Temperature | Cl.13 & 16 of IS 302-2-30 | HV Tester Line Testing | OEGPL | Available
Thermohygro Meter | 40 | 0.1⁰C, 1%RH,-50 - 70⁰C,10 - 99%RH | 19.05.22 |  |  | Lab ambient & Humidity | OEGPL | Required
Measuring Cylinder | 41 | 1ml, 0-100 ml | 19.05.22 | Resistance to rusting | Cl. 31 of IS 302-2-30 |  | OEGPL | Required
Ammonium Chloride | 42 |  |  | Resistance to rusting | Cl. 31 of IS 302-2-30 |  | OEGPL | NA
Carbon Tetrachloride | 43 |  |  | Resistance to rusting | Cl. 31 of IS 302-2-30 |  | OEGPL | NA